## VMS
AttrGrp | Attribute name | Attribute ID | Attribute type | Remark | Qualifier
| Handle | MDC_ATTR_ID_ HANDLE | HANDLE | Name binding attribute. | M
| System-Type | MDC_ATTR_SYS_ TYPE | TYPE | Examples: ventilator, monitor as defined in nomenclature. | M
| System-Model | MDC_ATTR_ID_ MODEL | SystemModel | Model describes manufacturer and model number. | C
| System-Id | MDC_ATTR_SYS_ID | OCTET STRING | Unique system ID, e.g., serial number. | C
| Compatibility-Id | MDC_ATTR_ID_ COMPAT | INT-U32 | For manufacturer use. | O
| Nomenclature- Version | MDC_ATTR_NOM_ VERS | Nomenclature-Version | Version of nomenclature used by the system. | C
| System- Capability | MDC_ATTR_SYS_ CAPAB | SystemCapability | Set of supported features; system specific. | O
| System- Specification | MDC_ATTR_SYS_ SPECN | SystemSpec | Defines functional components. | O
| Production- Specification | MDC_ATTR_ID_ PROD_SPECN | ProductionSpec | Component revisions, serial numbers, etc. | O
| Ext-Obj- Relations | MDC_ATTR_EXT_ OBJ_RELATION | ExtObjRelation- List | Relations to objects that are not defined in the DIM. | O

## MDS
AttrGrp | Attribute name | Attribute ID | Attribute type | Remark (section references are to 11073-10201) | Qualifier | PCD-01 Mapping | FHIR mapping | FHIR notes
| Handle | MDC_ATTR_ID_ HANDLE | HANDLE | Name binding attribute. | M | NM, Note A
| System-Type | MDC_ATTR_SYS_ TYPE | TYPE | Examples: ventilator, monitor as defined in nomenclature. | M
| System-Model | MDC_ATTR_ID_ MODEL | SystemModel | Model describes manufacturer and model number. | C | OBX under MDS
| System-Id | MDC_ATTR_SYS_ID | OCTET STRING | Unique system ID, e.g., serial number. | C | NM, Note A
| Compatibility-Id | MDC_ATTR_ID_ COMPAT | INT-U32 | For manufacturer use. | O | NM, Note A
| Nomenclature-Version | MDC_ATTR_NOM_ VERS | Nomenclature-Version | Version of nomenclature used by the system. | C | NM, Note A
| System-Capability | MDC_ATTR_SYS_ CAPAB | SystemCapability | Set of supported features; system specific. | O | NM, Note A
| System-Specification | MDC_ATTR_SYS_ SPECN | SystemSpec | Defines functional components. | O
| Production-Specification | MDC_ATTR_ID_ PROD_SPECN | ProductionSpec | Component revisions, serial numbers, etc. | O
| Ext-Obj-Relations | MDC_ATTR_EXT_ OBJ_RELATION | ExtObjRelation- List | Relations to objects that are not defined in the DIM. | O
| Mds-Status | MDC_ATTR_VMS_ MDS_STAT | MDSStatus | Device state according to MDS FSM. | C | OBX under MDS
| Bed-Label | MDC_ATTR_ID_BED_LABEL | OCTET STRING | Printable string identifying system location. | O | PV1-
| Soft-Id | MDC_ATTR_ID_SOFT | OCTET STRING | Settable, e.g., hospital inventory number. | O | OBX under MDS
| Operating-Mode | MDC_ATTR_MODE_ OP | PrivateOid |  | O | OBX under MDS
| Application-Area | MDC_ATTR_AREA_ APPL | ApplicationArea |  | O
| Patient-Type | MDC_ATTR_PT_ TYPE | PatientType | May control algorithms, see 7.10.1.1 for definition of type. | O
| Date-and-Time | MDC_ATTR_TIME_ ABS | AbsoluteTime | MDS maintains device time. | O
| Relative-Time | MDC_ATTR_TIME_ REL | RelativeTime |  | O
| HiRes-Relative- Time | MDC_ATTR_TIME_ REL_HI_RES | HighResRelativeTime |  | O
| Power-Status | MDC_ATTR_POWER_ STAT | PowerStatus | Either onBattery or onMains. | O
| Altitude | MDC_ATTR_ ALTITUDE | INT-I16 | Meters above or below sea level. | O
| Battery-Level | MDC_ATTR_VAL_ BAT_CHARGE | INT-U16 | In % of capacity; undefined if value > 100. | O
| Remaining- Battery-Time | MDC_ATTR_TIME_ BATT_REMAIN | BatMeasure | See 7.5.9.1 for the definition of type; minutes are the recommended measurement unit. | O
| Line-Frequency | MDC_ATTR_LINE_ FREQ | LineFrequency | Frequency of mains; implicitly in hertz (typically either 50 Hz or 60 Hz) | O | NM, Note A
| Association-Invoke-Id | MDC_ATTR_ID_ ASSOC_NO | INT-U16 | Counter for number of associations on a given communications port; incremented with each association control service element (ACSE) association | O | NM, Note A
| Locale | MDC_ATTR_LOCALE | Locale | Defines charset and language of printable string attributes in this MDS and contained objects. Contained MDS or VMD objects may define different Locale attributes for their scope. The top-level MDS shall support this attribute. | C

## MDS Notes
Note ref | Note
A | Not mapped in this revision. Either not needed because represented in a different way (e.g. by a relation in FHIR), or omitted because "80-20" rule not yet established. In PCD-01, may be included as OBX under MDS at option of sender (consumer may ignore)

## VMD
Attribute name | Attribute ID | Attribute type | Remark | Qualifier | FHIR maps as | FHIR notes
Type | MDC_ATTR_ID_TYPE | TYPE | Static type of this object; from object-oriented or metric nomenclature partition. Inherited from VMO | M
Handle | MDC_ATTR_ID_HANDLE | HANDLE | Locally unique short-hand identification. Inherited from VMO | M
VMD-Status | MDC_ATTR_VMD_STAT | VMDStatus | Example: on. | M
VMD-Model | MDC_ATTR_ID_MODEL | SystemModel | Manufacturer and model number. | O
Instance-Number | MDC_ATTR_ID_INSTNO | InstNumber | If multiple instances of the same VMD exist, this attribute helps to order the sequence. | O
Production-Specification | MDC_ATTR_ID_PROD_SPECN | ProductionSpec | Serial numbers and revisions; only present if VMD represents an independent subsystem. | O
Compatibility-Id | MDC_ATTR_ID_COMPAT | INT-U32 | Static for manufacturer use. | O
Parameter-Group | MDC_ATTR_ID_PARAM_GRP | OID-Type | Example: cardiovascular. | O
Position | MDC_ATTR_ID_POSN | INT-U16 | Example: slot number 0xffff marks an invalid or unknown position. | O
Operating-Hours | MDC_ATTR_TIME_PD_OP_HRS | INT-U32 |  | O
Operation-Cycles | MDC_ATTR_CYC_OP | INT-U32 | Example: number of measurements taken. | O
Measurement-Principle | MDC_ATTR_MSMT_PRINCIPLE | MsmtPrinciple | Describes the physical principle of the measurement. | O
Locale | MDC_ATTR_LOCALE | Locale | Defines charset and language | O

## Channel
Attribute name | Attribute ID | Attribute type | Remark | Qualifier | FHIR maps as | FHIR Notes
Type | MDC_ATTR_ID_TYPE | TYPE | Static type of this object; from object-oriented or metric nomenclature partition. Inherited from VMO | M
Handle | MDC_ATTR_ID_HANDLE | HANDLE | Locally unique short-hand identification. Inherited from VMO | M
Physical-Channel-No | MDC_ATTR_CHAN_ NUM_PHYS | INT-U16 | Provides a reference to a particular hardware channel, e.g., A/D. | O
Logical-Channel-No | MDC_ATTR_CHAN_ NUM_LOGICAL | INT-U16 | Dynamic channel numbering. | O
Parameter-Group | MDC_ATTR_ID_ PARAM_GRP | OID-Type | Static group of metrics, e.g., cardiovascular. | O
Measurement-Principle | MDC_ATTR_MSMT_ PRINCIPLE | MsmtPrinciple | Describes the physical principle of the measurement. | O
Color | MDC_ATTR_COLOR | SimpleColour | Useful to assign a common color to objects in one channel. | O

## Metric
Attribute name | Attribute ID | Attribute type | Remark | Qualifier | FHIR maps as | FHIR notes
Type | MDC_ATTR_ID_TYPE | TYPE | Static type of this object; from object-oriented or metric nomenclature partition. Inherited from VMO | M
Handle | MDC_ATTR_ID_HANDLE | HANDLE | Locally unique short-hand identification. Inherited from VMO | M
Metric-Status | MDC_ATTR_METRIC_STAT | MetricStatus |  | O
Measurement- Status | MDC_ATTR_MSMT_ STAT | MeasurementStatus | Usually part of an observed value. | O
Metric-Id-Partition | MDC_ATTR_METRIC_ID_PART | NomPartition | Identifies the nomenclature partition associated with the MetricId if it is different from the partition specified in the object’s VMO::Type attribute. | O
Metric-Id | MDC_ATTR_ID_ PHYSIO | OID-Type | Contains dynamic identification (e.g., a specific blood pressure label) compared to the static, generic ID in the Metric-Specification. OID is from VMO::Type or Metric-Id-Partition partition. Usually this attribute is part of an observed value, not an individual attribute. | O
Metric-Id-Ext | MDC_ATTR_ID_ MSMT_EXT | ExtNomenRef | Dynamic identification of the metric in a different nomenclature or dictionary. Use of this attribute severely limits interoperability of applications. | O
Unit-Code | MDC_ATTR_UNIT_ CODE | OID-Type | Example: mmHG; usually part of observed value. | O
Unit-LabelString | MDC_ATTR_UNIT_ LABEL_STRING | OCTET STRING | Textual representation of dimension. | O
Vmo-Source-List | MDC_ATTR_VMO_ LIST_SRC | VmoSourceList | Indicates sources of this metric in the form of references to other metrics. | O
Metric-Source- List | MDC_ATTR_METRIC_LIST_SRC | MetricSourceList | Indicates sources of this metric in the form of a list of metric IDs. | O
Msmt-Site-List | MDC_ATTR_SITE_ LIST_MSMT MDC_ATTR_SITE_ LIST_MSMT_EXT | SiteList SiteListExt | Measurement sites, specified in internal or external nomenclature. | O
Body-Site-List | MDC_ATTR_SITE_ LIST_BODY MDC_ATTR_SITE_ LIST_BODY_EXT | SiteList SiteListExt | Body sites, specified in internal or external nomenclature. | O
Metric-Calibration | MDC_ATTR_METRIC_CALIB | MetricCalibration | Indicates type and last time of calibration. | O
Color | MDC_ATTR_COLOR | SimpleColour | Color for representation. | O
Measure-Mode | MDC_ATTR_MODE_ MSMT | PrivateOid | Manufacturer-specific measurement information. | O
Measure-Period | MDC_ATTR_TIME_ PD_MSMT | MetricMeasure | Measurement repetition time; not necessarily the same as update period. | O
Averaging- Period | MDC_ATTR_TIME_ PD_AVG | MetricMeasure | Time period used to average values, e.g., a metric for the average flow of last hour. | O
Start-Time | MDC_ATTR_TIME_ START | AbsoluteTime | Time when measurement activity was started, e.g., when infusion was started. | O
Stop-Time | MDC_ATTR_TIME_ STOP | AbsoluteTime | Time when measurement activity was stopped. | O
Metric-Info-LabelString | MDC_ATTR_METRIC_INFO_LABEL_ STRING | OCTET STRING | Textual attribute, e.g., to specify electrode displacements or other specific information about the measurement. | O
Substance | MDC_ATTR_ID_ SUBSTANCE | ExtNomenRef | Substance to which a metric pertains; expressed in nomenclature that is defined outside of this standard. | O
Substance-Label-String | MDC_ATTR_ID_ SUBSTANCE_LABEL_STRING | OCTET STRING | Textual attribute that identifies the substance. | O

## Numeric
AttrGrp | Attribute name | Attribute ID | Attribute type | Remark | Qualifier | FHIR mapping | FHIR notes
| Nu-Observed-Value | MDC_ATTR_NU_VAl_OBS | NuObsValue | Example: measurement value; should also contain validity information to be useful. | C
| Compound-NuObserved-Value | MDC_ATTR_NU_ CMPD_VAL_OBS | NuObsValueCmp | Used when multiple values are represented in a single Numeric object. (Structure is compound.) | C
| Absolute-Time-Stamp | MDC_ATTR_TIME_ STAMP_ABS | AbsoluteTime | Time of observation (timestamp). | O
| Relative-Time-Stamp | MDC_ATTR_TIME_ STAMP_REL | RelativeTime |  | O
| HiRes-Time-Stamp | MDC_ATTR_TIME_ STAMP_REL_HI_RES | HighResRelativeTime |  | O
| Nu-Measure-Range | MDC_ATTR_NU_ RANGE_MSMT | AbsoluteRange | Potential measurement range. | O
| Nu-Physiological-Range | MDC_ATTR_NU_ RANGE_PHYSIO | AbsoluteRange | Physiological reasonable range (note that this is not an alarming range). | O
| Nu-Measure-Resolution | MDC_ATTR_NU_ MSMT_RES | FLOAT-Type | Resolution of measurement; minimum difference between two observed values. | O
| Display-Resolution | MDC_ATTR_DISP_ RES | DispResolution | Used when different resolution is needed when value is displayed. | O
| Accuracy | MDC_ATTR_NU_ ACCUR_MSMT | FLOAT-Type | Maximum deviation of actual value from reported observed value (if it can be specified). | O

## Enumeration
Attribute name | Attribute ID | Attribute type | Remark | Qualifier | FHIR mapping | FHIR notes
Enum-Observed-Value | MDC_ATTR_VAL_ ENUM_OBS | EnumObsValue | Either Enum-Observed-Value or Compound-Enum-Observed-Value shall be supported in one object instance. | C
Compound-Enum-Observed-Value | MDC_ATTR_VAL_ ENUM_OBS_CMPD | EnumObsValueCmp | Either Enum-Observed-Value or Compound-Enum-Observed-Value shall be supported in one object instance. | C
Absolute-Time- Stamp | MDC_ATTR_TIME_ STAMP_ABS | AbsoluteTime |  | O
Relative-Time- Stamp | MDC_ATTR_TIME_ STAMP_REL | RelativeTime |  | O
HiRes-Time-Stamp | MDC_ATTR_TIME_ STAMP_REL_HI_RES | HighResRelativeTime | High-resolution timestamp. | O
Enum-Measure-Range | MDC_ATTR_ENUM_ RANGE_MSMT | EnumMsmtRange | List of supported observed value OIDs. Optional if the OID type (EnumVal::enum-obj-id) is used in the observed value; out of the scope of this standard otherwise. | C
Enum-Measure-Range-Bit-String | MDC_ATTR_ENUM_ RANGE_MSMT_BIT_ STRING | BITS-32 | List of supported observed value bits in the bit string data type. Optional if the bit string type (EnumVal::enum-bit-str) is used in the observed value; out of the scope of this standard otherwise. | C
Enum-Measure-Range-Labels | MDC_ATTR_ENUM_ RANGE_MSMT_ LABELS | EnumMsmtRange Labels | Associates text strings with specific enumeration values. | O

## Data Types
SystemModel ::= SEQUENCE {
manufacturer | OCTET STRING,
model-number | OCTET STRING
}
System-Capability attribute is a top-level specification of implemented functions; (the following is an example only)
SystemCapability ::= BITS-32 {
sc-multiple-context(0), | indicates that System uses multiple naming contexts
sc-dyn-configuration(1), | containment tree changes dynamically
sc-dyn-scanner-create(2), | System allows host to create Scanner objects dynamically
sc-auto-init-scan-list(3), | CfgScanner object supports automatic
=-- scan list initialization
sc-auto-updt-scan-list(4) | CfgScanner object supports automatic scan list update
}
System-specification attribute allows specific entries for System functional components
SystemSpec ::= SEQUENCE OF SystemSpecEntry
SystemSpecEntry ::= SEQUENCE { component-capab-id | PrivateOid,
component-spec | ANY DEFINED BY component-capab-id
}
Production-Specification attribute deals with serial numbers, part numbers, revisions, etc.; note that a device may have multiple components so the Production-specification attribute should be a printable string defining the component and the "number"
ProductionSpec ::= SEQUENCE OF ProdSpecEntry
ProdSpecEntry ::= SEQUENCE {
spec-type INT-U16 {
unspecified(0),
serial-number(1),
part-number(2),
hw-revision(3),
sw-revision(4),
fw-revision(5),
protocol-revision(6),
prod-spec-gmdn(7) | Global Medical Device Nomenclature
},
component-id | PrivateOid,
prod-spec | OCTET STRING
}
NomenclatureVersion ::= SEQUENCE {
nom-major-version BITS-16 { | major version identifier
majorVersion1(0),
majorVersion2(1),
majorVersion3(2),
majorVersion4(3)
},
nom-minor-version INT-U16 | counter to identify minor updates
}
Note on A7: System-Capability attribute is a top-level specification of implemented functions; (the following is an example only)

## FHIR Location
Device.identifier.use
Device.identifier.system
Device.identifier.value
Device.identifier.assigner.reference
Device.manufacturer
Device.model
Device.type.coding.system
Device.type.coding.code
Device.type.coding.display
DeviceComponent.productionSpecification.specType.coding.system
DeviceComponent.productionSpecification.specType.coding.code
DeviceComponent.productionSpecification.specType.coding.display
DeviceComponent.productionSpecification.productionSpec
DeviceComponent.productionSpecification.componentId.use
DeviceComponent.productionSpecification.componentId.coding.text
DeviceComponent.productionSpecification.componentId.system
DeviceComponent.productionSpecification.componentId.assigner.reference
Observation.code.coding
Observation.value
Observation.valueQuantity.unit
Observation.effectiveTime
Observation.dataAbsentReason

## FHIR Resources
Device
<Device xmlns="http://hl7.org/fhir"> doco
<!-- from Resource: id, meta, implicitRules, and language -->
<!-- from DomainResource: text, contained, extension, and modifierExtension -->
<identifier><!-- 0..* Identifier Instance id from manufacturer, owner, and others --></identifier>
<type><!-- 1..1 CodeableConcept What kind of device this is --></type>
<note><!-- 0..* Annotation Device notes and comments --></note>
<status value="[code]"/><!-- 0..1 available ¦ not-available ¦ entered-in-error -->
<manufacturer value="[string]"/><!-- 0..1 Name of device manufacturer -->
<model value="[string]"/><!-- 0..1 Model id assigned by the manufacturer -->
<version value="[string]"/><!-- 0..1 Version number (i.e. software) -->
<manufactureDate value="[dateTime]"/><!-- 0..1 Manufacture date -->
<expiry value="[dateTime]"/><!-- 0..1 Date and time of expiry of this device (if applicable) -->
<udi value="[string]"/><!-- 0..1 FDA mandated Unique Device Identifier -->
<lotNumber value="[string]"/><!-- 0..1 Lot number of manufacture -->
<owner><!-- 0..1 Reference(Organization) Organization responsible for device --></owner>
<location><!-- 0..1 Reference(Location) Where the resource is found --></location>
<patient><!-- 0..1 Reference(Patient) If the resource is affixed to a person --></patient>
<contact><!-- 0..* ContactPoint Details for human/organization for support --></contact>
<url value="[uri]"/><!-- 0..1 Network address to contact device -->
</Device>
DeviceComponent
<DeviceComponent xmlns="http://hl7.org/fhir"> doco
<!-- from Resource: id, meta, implicitRules, and language -->
<!-- from DomainResource: text, contained, extension, and modifierExtension -->
<type><!-- 1..1 CodeableConcept What kind of component it is --></type>
<identifier><!-- 1..1 Identifier Instance id assigned by the software stack --></identifier>
<lastSystemChange value="[instant]"/><!-- 1..1 Recent system change timestamp -->
<source><!-- 0..1 Reference(Device) A source device of this component --></source>
<parent><!-- 0..1 Reference(DeviceComponent) Parent resource link --></parent>
<operationalStatus><!-- 0..* CodeableConcept Component operational status --></operationalStatus>
<parameterGroup><!-- 0..1 CodeableConcept Current supported parameter group --></parameterGroup>
<measurementPrinciple value="[code]"/><!-- 0..1 other ¦ chemical ¦ electrical ¦ impedance ¦ nuclear ¦ optical ¦ thermal ¦ biological ¦ mechanical ¦ acoustical ¦ manual+ -->
<productionSpecification> <!-- 0..* Production specification of the component -->
<specType><!-- 0..1 CodeableConcept Specification type --></specType>
<componentId><!-- 0..1 Identifier Internal component unique identification --></componentId>
<productionSpec value="[string]"/><!-- 0..1 A printable string defining the component -->
</productionSpecification>
<languageCode><!-- 0..1 CodeableConcept Language code for the human-readable text strings produced by the device --></languageCode>
</DeviceComponent>
DeviceMetric
<DeviceMetric xmlns="http://hl7.org/fhir"> doco
<!-- from Resource: id, meta, implicitRules, and language -->
<!-- from DomainResource: text, contained, extension, and modifierExtension -->
<type><!-- 1..1 CodeableConcept Type of metric --></type>
<identifier><!-- 1..1 Identifier Unique identifier of this DeviceMetric --></identifier>
<unit><!-- 0..1 CodeableConcept Unit of metric --></unit>
<source><!-- 0..1 Reference(Device) Describes the link to the source Device --></source>
<parent><!-- 0..1 Reference(DeviceComponent) Describes the link to the parent DeviceComponent --></parent>
<operationalStatus value="[code]"/><!-- 0..1 on ¦ off ¦ standby -->
<color value="[code]"/><!-- 0..1 black ¦ red ¦ green ¦ yellow ¦ blue ¦ magenta ¦ cyan ¦ white -->
<category value="[code]"/><!-- 1..1 measurement ¦ setting ¦ calculation ¦ unspecified -->
<measurementPeriod><!-- 0..1 Timing Describes the measurement repetition time --></measurementPeriod>
<calibration> <!-- 0..* Describes the calibrations that have been performed or that are required to be performed -->
<type value="[code]"/><!-- 0..1 unspecified ¦ offset ¦ gain ¦ two-point -->
<state value="[code]"/><!-- 0..1 not-calibrated ¦ calibration-required ¦ calibrated ¦ unspecified -->
<time value="[instant]"/><!-- 0..1 Describes the time last calibration has been performed -->
</calibration>
</DeviceMetric>

## FHIR Patient notes
Element | Card. | Type | Is Modifier | Summary | Short Label | Definition | Requirements | Comments
Patient.managingOrganization | 0..1 | Reference(Organization) |  | y | Organization that is the custodian of the patient record | Organization that is the custodian of the patient record | Need to know who recognizes this patient record, manages and updates it | There is only one managing organization for a specific patient record. Other organizations will have their own Patient record, and may use the Link property to join the records together (or a Person resource which can include confidence ratings for the association)
Patient.link | 0..* |  | y | y | Link to another patient resource that concerns the same actual person | Link to another patient resource that concerns the same actual patient | There are multiple usecases: * Duplicate patient records due to the clerical errors associated with the difficulties of identifying humans consistently, and * Distribution of patient information across multiple servers | There is no assumption that linked patient records have mutual links.
Patient.link.other | 1..1 | Reference(Patient¦RelatedPerson) | y | y | The other patient or related person resource that the link refers to | The other patient resource that the link refers to |  | Referencing a RelatedPerson here removes the need to use a Person record to associate a Patient and RelatedPerson as the same individual.
Patient.link.type | 1..1 | code | y | y | replace ¦ refer ¦ seealso - type of link | The type of link between this patient resource and another patient resource.
Note on A1: Format is ResourceName.componentName.componentName
Note on B1: Min and max repetitions n..m
Note on C1: Data type, resource reference, profiles, aggregations, class names, type references, etc. - see the wiki
Note on D1: If Y, indicates this element modifies the meaning of sibling or descendant elements. Defaults to N
Note on E1: If Y, element is included in the summary view (query with _summary=true). Default is N
Note on F1: Short description or code1 ¦ code2 ¦ code3 +
Note on G1: Full definition - required
Note on H1: Why element/resource is needed or why selected constraints chosen
Note on I1: Additional notes for implementers

## FHIR Data types
Coding
<[name] xmlns="http://hl7.org/fhir">
<!-- from Element: extension -->
<system value="[uri]"/><!-- 0..1 Identity of the terminology system -->
<version value="[string]"/><!-- 0..1 Version of the system - if relevant -->
<code value="[code]"/><!-- 0..1 Symbol in syntax defined by the system -->
<display value="[string]"/><!-- 0..1 Representation defined by the system -->
<userSelected value="[boolean]"/><!-- 0..1 If this coding was chosen directly by the user -->
</[name]>
Identifier
<[name] xmlns="http://hl7.org/fhir">
<!-- from Element: extension -->
<use value="[code]"/><!-- 0..1 usual ¦ official ¦ temp ¦ secondary (If known) -->
<type><!-- 0..1 CodeableConcept Description of identifier --></type>
<system value="[uri]"/><!-- 0..1 The namespace for the identifier -->
<value value="[string]"/><!-- 0..1 The value that is unique -->
<period><!-- 0..1 Period Time period when id is/was valid for use --></period>
<assigner><!-- 0..1 Reference(Organization) Organization that issued id (may be just text) --></assigner>
</[name]>
Quantity | SimpleQuantity lacks comparator
<[name] xmlns="http://hl7.org/fhir"> doco
<!-- from Element: extension -->
<value value="[decimal]"/><!-- 0..1 Numerical value (with implicit precision) -->
<comparator value="[code]"/><!-- 0..1 < ¦ <= ¦ >= ¦ > - how to understand the value -->
<unit value="[string]"/><!-- 0..1 Unit representation -->
<system value="[uri]"/><!-- ?? 0..1 System that defines coded unit form -->
<code value="[code]"/><!-- 0..1 Coded form of the unit -->
</[name]>
CodeableConcept
<[name] xmlns="http://hl7.org/fhir"> doco
<!-- from Element: extension -->
<coding><!-- 0..* Coding Code defined by a terminology system --></coding>
<text value="[string]"/><!-- 0..1 Plain text representation of the concept -->
</[name]>
code
code | string | Indicates that the value is taken from a set of controlled strings defined elsewhere (see Using codes for further discussion). Technically, a code is restricted to a string which has at least one character and no leading or trailing whitespace, and where there is no whitespace other than single spaces in the contents
regex: [^\s]+([\s]+[^\s]+)*
dateTime
Primitive Types
FHIR Name | Value Domain | XML Representation | JSON representation
boolean | true ¦ false | xs:boolean, except that 0 and 1 are not valid values | JSON boolean (true or false)
integer | A signed 32-bit integer (for larger values, use decimal) | xs:int, except that leading 0 digits are not allowed | JSON number (with no decimal point)
string | A sequence of Unicode characters | xs:string | JSON String
Note that strings SHALL NOT exceed 1MB in size
decimal | Rational numbers that have a decimal representation. See below about the precision of the number | xs:decimal, except that decimals may not use exponents, and leading 0 digits are not allowed | A JSON number, except that exponents are not allowed
uri | A Uniform Resource Identifier Reference (RFC 3986 ). Note: URIs are case sensitive. For UUID (urn:uuid:53fefa32-fcbb-4ff8-8a92-55ee120877b7) use all lowercase | xs:anyURI | A JSON string - a URI
URIs can be absolute or relative, and may have an optional fragment identifier
base64Binary | A stream of bytes, base64 encoded (RFC 4648 ) | xs:base64Binary | A JSON string - base64 content
Todo: is it possible to impose an upper absolute limit on a base64Binary (for denial of service reasons, like on string)
instant | An instant in time - known at least to the second and always includes a time zone. Note: This is intended for precisely observed times (typically system logs etc.), and not human-reported times - for them, use date and dateTime. instant is a more constrained dateTime | xs:dateTime | A JSON string - an xs:dateTime
Note: This type is for system times, not human times (see date and dateTime below).
date | A date, or partial date (e.g. just year or year + month) as used in human communication. There is no time zone. Dates SHALL be valid dates | union of xs:date, xs:gYearMonth, xs:gYear | A JSON string - a union of xs:date, xs:gYearMonth, xs:gYear
Regex: -?[0-9]{4}(-(0[1-9]¦1[0-2])(-(0[0-9]¦[1-2][0-9]¦3[0-1]))?)?
dateTime | A date, date-time or partial date (e.g. just year or year + month) as used in human communication. If hours and minutes are specified, a time zone SHALL be populated. Seconds must be provided due to schema type constraints but may be zero-filled and may be ignored. Dates SHALL be valid dates. The time "24:00" is not allowed | union of xs:dateTime, xs:date, xs:gYearMonth, xs:gYear | A JSON string - a union of xs:dateTime, xs:date, xs:gYearMonth, xs:gYear
Regex: -?[0-9]{4}(-(0[1-9]¦1[0-2])(-(0[0-9]¦[1-2][0-9]¦3[0-1])(T([01][0-9]¦2[0-3]):[0-5][0-9]:[0-5][0-9](\.[0-9]+)?(Z¦(\+¦-)((0[0-9]¦1[0-3]):[0-5][0-9]¦14:00)))?)?)?
Annotation
<[name] xmlns="http://hl7.org/fhir">
<!-- from Element: extension -->
<author[x]><!-- 0..1 Reference(Practitioner¦Patient¦RelatedPerson)¦string Individual responsible for the annotation --></author[x]>
<time value="[dateTime]"/><!-- 0..1 When the annotation was made -->
<text value="[string]"/><!-- 1..1 The annotation - text content -->
</[name]>
Reference
<[name] xmlns="http://hl7.org/fhir"> doco
<!-- from Element: extension -->
<reference value="[string]"/><!-- ?? 0..1 Relative, internal or absolute URL reference -->
<display value="[string]"/><!-- 0..1 Text alternative for the resource -->
</[name]>